## Main
Cluster | Partner | technique | probed area | probed area resolution | penetration depth | pernetration depth resolution | detection limit | detection limit Resolution | contrast | observables | Observables resolution | characterization mode | time scale | Sample preparation | Instrument | Beam source | Cell Availiability | Degree of fidelity | Maturity | Resources | Planning time | Post-processing time | CRL by observable Acquisition | CRL by observable analysis | used in BIG-map | Sample reuse | Data analysis Software
[11 empty] | High/Medium/Low | Ex-situ/In-situ/Operando |  |  |  |  | Yes/no | Low/medium/high/best | Commercial/Realistic/Model | Cheap/Medium/Expensive | Hours/Days/Weeks/Months | Hours/Days/Weeks/Months | High-Medium-Low | High-Medium-Low | Yes/No | Yes/No | Inhouse/commercial/opensource/name
Diffraction | [DTU] | [X-ray diffraction, XRD] | [few mm2] | [Null] | [mm, cm] | [Null] | [crystallinity (> 5 nm cryst. size)] | [1-2 wt.%] | [atomic number contrast] | [crystal structure, phase fractions, crystallite morphology, strain] | [Null] | [Ex-situ, operando] | [ms, min] | [Cylindrical cell, Pouch cell] | [Null] | [Synchrotron] |  | [Best] | [Commercial] | [Expensive] | [Months] | [Hours/Days] | [High] | [Medium] | [Yes] | [Yes] | [Commercial (Topas), free (Fullprof)]
Diffraction | [ESRF] | [X-ray diffraction, XRD] | [few mm2] | [Null] | [cm] | [Null] | [crystallinity (> 5 nm cryst. size)] | [1-2 wt.%] | [atomic number contrast] | [crystal structure, phase fractions, crystallite morphology, strain] | [Null] | [Ex-situ, operando] | [ms, min] | [transmission geometry, reflection geometry, CR20XX coin cell, swagelock type cell] | [ID22, ID31] | [Synchrotron] |  | [Best] | [Commercial] | [Expensive] | [Months] | [Hours/Days]
Diffraction | [CEA] | [Neutron Diffraction, ND] | [up to 3x5 cm2] | [5 mm] | [cm] | [Null] | [crystallinity (> 5 nm cryst. size)] | [1-2 wt.%] | [Neutron scattering contrast] | [crystal structure, phase fractions, crystallite morphology, strain] | [H,H,L,M] | [Ex-situ, operando] | [min] | [Model Cell, Cylindrical cell, Pouch cell] | [D20, D19] | [Neutron] | [Yes] | [High] | [Commercial] | [Expensive] | [Months] | [Hours/Days] | [High, High, Medium, High] | [High, High, Medium, High] | [yes] | [yes] | [FullProf]
Diffraction | [ILL] | [Neutron Diffraction, ND] | [up to 3x5 cm2] | [1 mm] | [cm] | [Null] | [crystallinity (> 5 nm cryst. size)] | [1-2 wt.%] | [Neutron scattering contrast] | [crystal structure, phase fractions, crystallite morphology, strain] | [High] | [Ex-situ, operando] | [ms, min] | [ILL operando cells] | [D2B, D20, D19] | [Neutron] | [Yes] | [High/best] | [Commercial] | [Expensive] | [Months] | [Hours/Days] | [High, High, Medium, High] | [High, High, Medium, High] | [Yes] | [Yes] | [FullProf]
Diffraction | [TUD] | [Neutron Diffraction, ND] | [up to 3x5 cm2] | [1 mm] | [cm] | [Null] | [crystallinity (> 5 nm cryst. size)] | [1-2 wt.%] | [Neutron scattering contrast] | [crystal structure, phase fractions, crystallite morphology, strain] | [Null] | [Ex-situ, operando] | [ms, min] | [Null] | [PEARL] | [Neutron] |  | [High/best] | [Commercial] | [Expensive] | [Months] | [Hours/Days]
Diffraction | [TUD] | [Laboratory X-ray Diffraction, lab XRD] | [1 to 50 mm2] | [Null] | [101 μm with Cu (8 keV), 1 mm with Mo (17.45 keV)] | [Null] | [crystallinity (> 5 nm cryst. size)] | [1-2 wt.%] | [atomic number contrast] | [crystal structure, phase fractions, crystallite morphology, strain] | [Null] | [Ex-situ, operando] | [5 < Δt < 30 mins] | [Kapton-window cell, gas connections, reflection geometry] | [Null] | [Cu, Co, Mo, Ag] |  | [High] | [realistic] | [Cheap] | [Hours] | [Hours/Days]
Diffraction | [CEA] | [Laboratory X-ray Diffraction, lab XRD] | [1 to 50 mm2] | [cm] | [101 μm with Cu (8 keV), 1 mm with Mo (17.45 keV)] | [Null] | [crystallinity (> 5 nm cryst. size)] | [1-2 wt.%] | [atomic number contrast] | [crystal structure, phase fractions, crystallite morphology, strain] | [M, H, L, M] | [Ex-situ, operando] | [5 < Δt < 30 mins] | [Be windows, coin cells] | [Bruker D8 Advance] | [Cu, Mo] | [Yes] | [High] | [realistic] | [Cheap] | [Hours] | [Hours/Days] | [High, High, Medium, High] | [High, High, Medium, High] | [no] | [yes] | [Fullprof]
Diffraction | [DTU] | [Laboratory X-ray Diffraction, lab XRD] | [1 to 50 mm2] | [Null] | [101 μm with Cu (8 keV)] | [Null] | [crystallinity (> 5 nm cryst. size)] | [1-2 wt.%] | [atomic number contrast] | [crystal structure, phase fractions, crystallite morphology, strain] | [Null] | [Ex-situ, operando] | [5 < Δt < 30 mins] | [ELCELL, reflection geometry] | [Null] | [Cu] | [Yes] | [High] | [realistic] | [Cheap] | [Hours] | [Hours/Days] | [High] | [High] | [Yes] | [No] | [Commercial (Topas), free (Fullprof)]
Diffraction | [DTU] | [Laboratory X-ray Diffraction, lab XRD] | [1 to 50 mm2] | [Null] | [101 μm with Cu (8 keV)] | [Null] | [crystallinity (> 5 nm cryst. size)] | [1-2 wt.%] | [atomic number contrast] | [crystal structure, phase fractions, crystallite morphology, strain] | [Null] | [Ex-situ, operando] | [5 < Δt < 30 mins] | [AMPIX cell, pouch cell, transmission geometry] | [Null] | [Cu] | [Yes] | [High] | [realistic] | [Cheap] | [Hours] | [Hours/Days] | [High] | [High] | [Yes] | [Yes] | [Commercial (Topas), free (Fullprof)]
Diffraction | [CNRS] | [Laboratory X-ray Diffraction, lab XRD] | [1 to 50 mm2] | [Null] | [101 μm with Cu (8 keV), 1 mm with Mo (17.45 keV)] | [Null] | [crystallinity (> 5 nm cryst. size)] | [1-2 wt.%] | [atomic number contrast] | [crystal structure, phase fractions, crystallite morphology, strain] | [Null] | [Ex-situ, operando] | [5 < Δt < 30 mins] | [Glassy carbon window cell, reflection géometry] | [Null] | [Cu, Co, Mo] |  | [High] | [realistic] | [Cheap] | [Hours] | [Hours/Days]
Diffraction | [CNRS] | [Laboratory X-ray Diffraction, lab XRD] | [1 to 50 mm2] | [Null] | [101 μm with Cu (8 keV), 1 mm with Mo (17.45 keV)] | [Null] | [crystallinity (> 5 nm cryst. size)] | [1-2 wt.%] | [atomic number contrast] | [crystal structure, phase fractions, crystallite morphology, strain] | [Null] | [Ex-situ, operando] | [5 < Δt < 30 mins] | [Kapton & glassy carbon windows cell, transmition geometry] | [Null] | [Cu, Co, Mo] |  | [High] | [realistic] | [Cheap] | [Hours] | [Hours/Days]
Diffraction | [CNRS] (LRCS) | [Laboratory X-ray Diffraction, lab XRD] | [1 to 50 mm2] | [Null] | [101 μm with Cu (8 keV), 1 mm with Mo (17.45 keV)] | [Null] | [crystallinity (> 5 nm cryst. size)] | [1-2 wt.%] | [atomic number contrast] | [crystal structure, phase fractions, crystallite morphology, strain] | [Null] | [Ex-situ, operando] | [5 < Δt < 30 mins] | [Be-window cell, refection geometry] | [Null] | [Cu, Co, Mo] |  | [High] | [realistic] | [Cheap] | [Hours] | [Hours/Days]
Diffraction | [CNRS] (LRCS) | [Laboratory X-ray Diffraction, lab XRD] | [1 to 50 mm2] | [Null] | [101 μm with Cu (8 keV), 1 mm with Mo (17.45 keV)] | [Null] | [crystallinity (> 5 nm cryst. size)] | [1-2 wt.%] | [atomic number contrast] | [crystal structure, phase fractions, crystallite morphology, strain] | [Null] | [Ex-situ, operando] | [5 < Δt < 30 mins] | [Be-window cell, refection geometry] | [Null] | [Cu, Co, Mo] |  | [High] | [realistic] | [Cheap] | [Hours] | [Hours/Days]
Diffraction | [UU] | [Laboratory X-ray Diffraction, lab XRD] | [1 to 50 mm2] | [Null] | [101 μm with Cu (8 keV), 1 mm with Mo (17.45 keV)] | [Null] | [crystallinity (> 5 nm cryst. size)] | [1-2 wt.%] | [atomic number contrast] | [crystal structure, phase fractions, crystallite morphology, strain] | [Null] | [Ex-situ, operando] | [5 < Δt < 30 mins] | [Pouch cell, transmission geometry] | [Null] | [Cu] |  | [High] | [realistic] | [Cheap] | [Hours] | [Hours/Days]
Diffraction | [Soleil] | [Synchrotron X-ray diffraction, synchrotron XRD] | [few mm2] | [Null] | [mm, cm] | [Null] | [crystallinity (> 5 nm cryst. size)] | [1-2 wt.%] | [atomic number contrast] | [crystal structure, phase fractions, crystallite morphology, strain] | [Null] | [Ex-situ, operando] | [1 s < Δt < 5 min] | [Be-window cell, reflection geometry] | [cristal, swing] | [Synchrotron] |  | [Best] | [realistic] | [Expensive] | [Months] | [Hours/Days]
Diffraction | [Soleil] | [Synchrotron X-ray diffraction, synchrotron XRD] | [few mm2] | [Null] | [mm, cm] | [Null] | [crystallinity (> 5 nm cryst. size)] | [1-2 wt.%] | [atomic number contrast] | [crystal structure, phase fractions, crystallite morphology, strain] | [Null] | [Ex-situ, operando] | [1 s < Δt < 5 min] | [Kapton window cell, transmission geometry] | [cristal, swing] | [Synchrotron] |  | [Best] | [realistic] | [Expensive] | [Months] | [Hours/Days]
Diffraction | [CNRS] | [Synchrotron X-ray diffraction, synchrotron XRD] | [few mm2] | [Null] | [mm, cm] | [Null] | [crystallinity (> 5 nm cryst. size)] | [1-2 wt.%] | [atomic number contrast] | [crystal structure, phase fractions, crystallite morphology, strain] | [Null] | [Ex-situ, operando] | [1 s < Δt < 5 min] | [Kapton & glassy carbon windows cell, transmition geometry] | [Null] | [Synchrotron] |  | [Best] | [realistic] | [Expensive] | [Months] | [Hours/Days]
Diffraction | [DTU] | [Synchrotron X-ray diffraction, synchrotron XRD] | [few mm2] | [Null] | [mm, cm] | [Null] | [crystallinity (> 5 nm cryst. size)] | [1-2 wt.%] | [atomic number contrast] | [crystal structure, phase fractions, crystallite morphology, strain] | [Null] | [Ex-situ, operando] | [1 s < Δt < 5 min] | [6-position AMPIX cells] | [DanMAX (MAX IV)] | [Synchrotron] | [Yes] | [Best] | [realistic] | [Expensive] | [Months] | [Hours/Days] | [High] | [High] | [Yes] | [Yes] | [Commercial (Topas), free (Fullprof)]
Diffraction | [UU] | [Synchrotron X-ray diffraction, synchrotron XRD] | [few mm2] | [Null] | [mm, cm] | [Null] | [crystallinity (> 5 nm cryst. size)] | [1-2 wt.%] | [atomic number contrast] | [crystal structure, phase fractions, crystallite morphology, strain] | [Null] | [Ex-situ, operando] | [1 s < Δt < 5 min] | [AMPIX-style cell, glassy carbon window, transmission geometry] | [Null] | [Synchrotron] |  | [Best] | [realistic] | [Expensive] | [Months] | [Hours/Days]
Diffraction | [UU] | [Synchrotron X-ray diffraction, synchrotron XRD] | [few mm2] | [Null] | [mm, cm] | [Null] | [crystallinity (> 5 nm cryst. size)] | [1-2 wt.%] | [atomic number contrast] | [crystal structure, phase fractions, crystallite morphology, strain] | [Null] | [Ex-situ, operando] | [1 s < Δt < 5 min] | [Pouch cell with glassy carbon windows, transmission geometry] | [Null] | [Synchrotron] |  | [Best] | [realistic] | [Expensive] | [Months] | [Hours/Days]
Diffraction | [TUD] | [Synchrotron X-ray diffraction, synchrotron XRD] | [few mm2] | [Null] | [mm, cm] | [Null] | [crystallinity (> 5 nm cryst. size)] | [1-2 wt.%] | [atomic number contrast] | [crystal structure, phase fractions, crystallite morphology, strain] | [Null] | [Ex-situ, operando] | [1 s < Δt < 5 min] | [Kapton-window cell, transmission geometry, gas connections] | [Null] | [Synchrotron] |  | [Best] | [realistic] | [Expensive] | [Months] | [Hours/Days]
Diffraction | [CEA] | [Synchrotron X-ray diffraction, synchrotron XRD] | [few mm2] | [5 - 50 μm] | [mm] | [Null] | [crystallinity (> 5 nm cryst. size)] | [0.5 wt.%] | [atomic number contrast] | [crystal structure, phase fractions, crystallite morphology, strain] | [H,H,L,M] | [Ex-situ, operando] | [1 s < Δt < 5 min] | [Pouch Cell, Small Swagelok Cell] | [BM02, ID31, ID13] | [Synchrotron] | [Yes] | [Best] | [realistic] | [Expensive] | [Months] | [Hours/Days] | [High, High, Medium, High] | [High, High, Medium, High] | [yes] | [yes] | [Python, PyFAI, Fullprof]
Diffraction | [ESRF] | [Synchrotron X-ray diffraction, synchrotron XRD] | [few mm2] | [Null] | [mm, cm] | [Null] | [crystallinity (> 5 nm cryst. size)] | [1-2 wt.%] | [atomic number contrast] | [crystal structure, phase fractions, crystallite morphology, strain] | [Null] | [Ex-situ, operando] | [1 s < Δt < 5 min] | [CR20XX cell, transmission geometry, reflection geometry] | [BM32, ID22, ID31, ID11, ID15A] | [Synchrotron] |  | [Best] | [realistic] | [Expensive] | [Months] | [Hours/Days]
Diffraction | [ESRF] | [Synchrotron X-ray diffraction, synchrotron XRD] | [few mm2] | [Null] | [mm, cm] | [Null] | [crystallinity (> 5 nm cryst. size)] | [1-2 wt.%] | [atomic number contrast] | [crystal structure, phase fractions, crystallite morphology, strain] | [Null] | [Ex-situ, operando] | [1 s < Δt < 5 min] | [Modified Swagelok-type cell] | [BM32, ID22, ID31, ID11, ID15A] | [Synchrotron] |  | [Best] | [realistic] | [Expensive] | [Months] | [Hours/Days]
Diffraction | [TUD] | [Synchrotron nano-diffraction] | [1 cm2] | [Null] | [Electrode surface] | [Null] | [crystallinity] | [Null] | [atomic number contrast] | [Crystal structure, strain] | [Null] | [Ex-situ, operando] | [< 1 hr] | [Kapton window dome cell, gas connections, reflection geometry] | [Null] | [Synchrotron] |  | [Medium/high] | [model] | [Expensive] | [Months] | [Hours/Days]
Diffraction | [ESRF] | [Synchrotron nano-diffraction] | [1 cm2] | [50 - 80 nm] | [Electrode surface] | [nm - µm] | [crystallinity] | [Null] | [atomic number contrast] | [Crystal structure, strain] | [H, H] | [Ex-situ, operando] | [<= 1 hr] | [windowed coin cell] | [ID11, ID01] | [Synchrotron] | [Yes] | [High] | [model] | [Expensive] | [Months] | [Hours/Days] | [Low] | [Medium/High] | [yes] | [Yes/No] | [Inhouse, Python]
SAXS | [DTU] | [laboratory Small Angle X-ray Scattering, lab SAXS] | [< 1 mm2] | [Null] | [micrometers] | [Null] | [0.1 wt.% ] | [Null] | [atomic weight contrast] | [Morphology (1-100 nm), pore structure, surface area] | [Null] | [Ex-situ, operando] | [Minutes] | [Capillary] | [Null] | [Xenocs] |  | [Low] | [realistic] | [Cheap] | [Days] | [Minutes/hours]
SAXS | [CTH] | [laboratory Small Angle X-ray Scattering, lab SAXS] | [< 1 mm2] | [Null] | [micrometers] | [Null] | [0.1 wt.% ] | [Null] | [atomic weight contrast] | [Morphology (1-100 nm), pore structure, surface area] | [Null] | [Ex-situ, operando] | [Minutes] | [Kapton window cell, diamond windows AMPIX cell, transmission geometry] | [Null] | [Mat:Nordic] | [Yes, but more testing is needed] | [Low] | [realistic] | [Medium] | [Weeks] | [Days] | [High] | [Medium] | [No] | [Yes] | [Matlab]
SAXS | [CEA] | [laboratory Small Angle X-ray Scattering, lab SAXS] | [< 1 mm2] | [< 1 mm2] | [100 μm - 1 mm] | [Null] | [1 wt.% ] | [Null] | [atomic weight contrast] | [Morphology (1-100 nm), pore structure, surface area] | [Null] | [Ex-situ] | [Minutes] | [Ex-situ cells] | [Cu, Mo] | [Xeuss 3.0] | [No] | [Low] | [realistic] | [Cheap] | [Days] | [Minutes/hours] |  |  | [No] | [Yes]
SAXS | [CEA] | [Synchrotron small angle X‑ray scattering, synchrotron ultra-small angle X-ray scattering, U-SAXS] | [1 cm2] | [< 1 mm2] | [< 10 mm] | [Null] | [0.1 wt.% ] | [Null] | [atomic weight contrast] | [Morphology, 1 nm to few μm, Pore Structure, Surface area] | [Null] | [Ex-situ, operando] | [> 1s] | [Pouch cell, SAXS-WAXS sample holder] | [Synchrotron] | [BM02, ID02, ID31] | [Yes] | [High] | [realistic] | [Expensive] | [Months] | [Minutes/hours] |  |  | [Yes] | [Yes] | [Python]
SAXS | [SOLEIL] | [Synchrotron small angle X‑ray scattering, synchrotron ultra-small angle X-ray scattering, U-SAXS] | [1 cm2] | [Null] | [cm] | [Null] | [0.1 wt.% ] | [Null] | [atomic weight contrast] | [Short range order, 10 nm to few μm, Pore Structure, Surface area] | [Null] | [Ex-situ, operando] | [> 1s] | [Null] | [Synchrotron] | [SWING] |  | [High] | [realistic] | [Expensive] | [Months] | [Minutes/hours]
SAXS | [ESRF] | [Synchrotron small angle X‑ray scattering, synchrotron ultra-small angle X-ray scattering, U-SAXS] | [1 cm2] | [Null] | [cm] | [Null] | [0.1 wt.% ] | [Null] | [atomic weight contrast] | [Short range order, 10 nm to few μm, Pore Structure, Surface area] | [Null] | [Ex-situ, operando] | [> 1s] | [Null] | [Synchrotron] | [BM02, ID02] |  | [High] | [realistic] | [Expensive] | [Months] | [Minutes/hours]
SAXS | [CEA] | [Synchrotron microbeam small angle scattering, Synchrotron microbeam SAXS, Synchrotron microbeam wide angle scattering, Synchrotron microbeam WAXS] | [1 cm2] | [400 nm] | [Electrode surface] | [Null] | [Null] | [Null] | [atomic weight contrast] | [morphology 1 to 100 nm, Pore Structure, Surface area] | [Null] | [Ex-situ, operando] | [< 1 min] | [Miniaturized cell, SAXS-WAXS sample holder] | [Synchrotron] | [ID31, ID13] |  | [Medium] | [realistic] | [Expensive] | [Months] | [Minutes/hours]
SAXS | [ESRF] | [Synchrotron microbeam small angle scattering, Synchrotron microbeam SAXS, Synchrotron microbeam wide angle scattering, Synchrotron microbeam WAXS] | [1 cm2] | [400 nm] | [Electrode surface] | [Null] | [Null] | [Null] | [atomic weight contrast] | [morphology 1 to 100 nm, Pore Structure, Surface area] | [Null] | [Ex-situ, operando] | [< 1 hr] | [Null] | [Synchrotron] | [ID13] |  | [Medium] | [realistic] | [Expensive] | [Months] | [Minutes/hours]
SAXS | [SOLEIL] | [Rheological Synchrotron small angle X‑ray scattering, Rheo-SAXS] | [1 cm2] | [Null] | [cm] | [Null] | [0.1 wt.% ] | [Null] | [atomic weight contrast] | [reological properties] | [Null] | Ex-situ, operando | [> 1s] | [Pouch cell, SAXS-WAXS sample holder] | [Synchrotron] | [SWING] |  |  | [realistic]
SANS | [CEA] | [Small Angle Neutron Scattering, SANS] | [55 x 40 mm2] | [10 to 300 mm2] | [Full cell] | [bulk] | [Null] | [Null] | [neutron scattering contrast] | [Morphology (1-100 nm), Size of nanoparticles, shape of nanoparticles, volume expansion, interface roughness] | [Null] | [Ex-situ, operando] | [2 min (regular beam), 15 min (pencil beam)] | [Ti-cell w/o separator] | [Neutron] | [D22] | [Null] | [Medium] | [realistic] | [Expensive] | [Months] | [Hours]
SANS | [ILL] | [Small Angle Neutron Scattering, SANS] | [55 x 40 mm2, 10 to 300 mm2] | [10 to 300 mm2] | [Full cell] | [bulk] | [Null] | [Null] | [neutron scattering contrast] | [Morphology (1-100 nm), Size of nanoparticles, shape of nanoparticles, volume expansion, interface roughness] | [Null] | [Ex-situ, operando] | [Null] | [Null] | [Neutron] | [D22 simultaneous SANS-SAXS] | [Null] | [Medium] | [realistic] | [Expensive] | [Months] | [Hours] |  |  |  |  | [Grasp]
SANS | [TUD] | [Small Angle Neutron Scattering, SANS] | [55 x 40 mm2, 10 to 300 mm2] | [10 to 300 mm2] | [Full cell] | [bulk] | [Null] | [Null] | [neutron scattering contrast] | [Morphology (1-100 nm), Size of nanoparticles, shape of nanoparticles, volume expansion, interface roughness] | [Null] | [Ex-situ, operando] | [Null] | [Null] | [Neutron] | [SANS-2] |  | [Medium] | [realistic] | [Expensive] | [Months] | [Hours]
SANS | [CEA] | [Two dimension Small angle neutron scattering, 2D SANS] | [1 mm2] | [Null] | [10-20 μm] | [bulk] | [Null] | [Null] | [neutron scattering contrast] | [Morphology (1-100 nm)] | [Null] | [Ex-situ, operando] | [Few minutes] | [Ti-PEEK based 2D cell (miniaturized, pencil beam)] | [Neutron] | [D22] | [Null] | [High] | [realistic] | [Expensive] | [Months] | [Hours]
SANS | [ILL] | [Two dimension Small angle neutron scattering, 2D SANS] | [1 mm2] | [Null] | [10-20 μm] | [bulk] | [Null] | [Null] | [neutron scattering contrast] | [Morphology (1-100 nm)] | [Null] | [Ex-situ, operando] | [Few minutes] | [Null] | [Neutron] | [D22] | [Null] | [High] | [realistic] | [Expensive] | [Months] | [Hours] | [High] | [High] | [No] | [No] | [Grasp]
Imaging | [CEA] | [X-ray micro-tomography, X-ray micro-CT] | [> 200 μm x 200 μm] | [< 1 μm] | [cm, full cell] | [ < 50 nm] | [atomic weight contrast] | [Null] | [Phase contrast boundary enhancement] | [Morphological information, volume variations, pre-existing defects, after ageing defects, failure mecahnisms] | [H, H, H, H, H] | [Ex-situ, operando] | [μs, s] | [cylindrical cells] | [Null] | [Synchrotron] | [Yes] | [High] | [Commercial] | [Expensive] | [Months] | [Hours/Days] | [High] | [Medium] | [Yes] | [Yes] | [Null]
Imaging | [ESRF] | [X-ray micro-tomography, X-ray micro-CT] | [> 200 μm x 200 μm] | [< 1 μm] | [cm, full cell] | [ < 50 nm] | [atomic weight contrast] | [Null] | [Phase contrast boundary enhancement] | [Morphological information, volume variations, pre-existing defects, after ageing defects, failure mecahnisms] | [H, H, H, H, H] | [Ex-situ, operando] | [μs, s] | [cell for imaging battery explosion] | [ID19, BM05] | [Synchrotron] | [No] | [High] | [Commercial] | [Expensive] | [Months] | [Hours/Days] | [High] | [Medium] | [Yes] | [Yes] | [Inhouse]
Imaging | [SOLEIL] | [X-ray micro-tomography, X-ray micro-CT] | [> 200 μm x 200 μm] | [< 1 μm] | [cm, full cell] | [ < 50 nm] | [atomic weight contrast] | [Null] | [Phase contrast boundary enhancement] | [Morphological information, volume variations, pre-existing defects, after ageing defects, failure mecahnisms] | [H, H, H, H, H] | [Ex-situ, operando] | [μs, s] | [Null] | [PSICHé] | [Synchrotron] | [No] | [High] | [Commercial] | [Expensive] | [Months] | [Hours/Days] | [High] | [Medium] | [No] | [Yes] | [Null]
Imaging | [ESRF] | [X-ray diffraction computed tomography, XRD-CT] | [1 cm x 1 cm] | [300 nm – 1 μm] | [cm, full cell] | [Null] | [crystallinity] | [Null] | [crystal structure contrast] | [Phase fraction (SOC), Unit cell lattice parameters, atomic site parameters, atomic thermal parameters , lattice strain, lattice size, structural information, chemical information, morphological information] | [Null, H, M, L, L, L, H, H, H] | [Ex-situ, operando] | [> 10 ms] | [Null] | [ID15A, ID06, ID11] | [Synchrotron] |  | [Medium] | [Commercial] | [Expensive] | [Months] | [Hours/Days]
Imaging | [ESRF] | [X-ray total scattering computed tomography, Pair Distribution Function computed tomography, X-ray PDF-CT] | [1 cm x 1 cm] | [< 1 μm] | [cm, full cell] | [Null] | [crystal structure] | [Null] | [local structure contrast] | [Structural information, chemical information, morphological information] | [H, H, H] | [Ex-situ, operando] | [> 10 ms] | [Null] | [ID15A] | [Synchrotron] |  | [Low] | [Commercial] | [Medium/Expensive] | [Months] | [Hours/Days]
Imaging | [CEA] | [neutron imaging] | [5 mm x 5 mm, 170 mm x 170 mm] | [2.5 μm x 2.5 μm, < 4 μm x 4 μm ] | [cm, fuel cell] | [Null] | [neutron absorption contrast, lithium distribution, isotopes labeling (e.g. 6Li, 7Li)] | [Null] | [possible isotopes labeling (e.g. 6Li, 7Li)] | [lithium distribution , lithium concentration gradient, volume variations, defects after ageing] | [H, M, M, M] | [Ex-situ, operando] | [s, min] | [coin cell, small size cell] | [NeXT] | [Neutron] | [Yes] | [High] | [Commercial] | [Expensive] | [Months] | [Hours/Days] | [High] | [Medium] | [Yes] | [Yes] | [Inhouse, Pyhton]
Imaging | [ILL] | [neutron imaging] | [> 4 μm x 4 μm] | [Null] | [cm, full cell] | [Null] | [neutron absorption contrast, isotopes labeling (e.g. 6Li, 7Li)] | [Null] | [possible isotopes labeling (e.g. 6Li, 7Li)] | [lithium distribution , lithium concentration gradient, volume variations, defects after ageing] | [L, L, L, Null] | [Ex-situ, operando] | [ms, s] | [operando cell, ex-situ cell] | [D50-Next] | [Neutron] | [No] | [High] | [Commercial] | [Expensive] | [Months] | [Hours] |  |  | [Yes] |  | [Python]
Imaging | [CEA] | [Neutron tomography] | [> 4 mm] | [10 μm] | [cm, full cell] | [Null] | [neutron absorption contrast, lithium distribution, isotopes labeling (e.g. 6Li, 7Li) ] | [Null] | [possible isotopes labeling (e.g. 6Li, 7Li)] | [lithium distribution , lithium concentration gradient, volume variations, defects after ageing] | [M, M, M, M] | [Ex-situ] | [hr] | [Coin cells, small sized cells] | [NeXT] | [Neutron] | [Yes] | [High] | [Commercial] | [Expensive] | [Months] | [Hours/Days] | [High] | [Medium] | [No] | [Yes] | [Ihouse, Python]
Imaging | [ILL] | [Neutron tomography] | [> 4 mm] | [Null] | [cm, full cell] | [Null] | [neutron absorption contrast, isotopes labeling (e.g. 6Li, 7Li)] | [Null] | [possible isotopes labeling (e.g. 6Li, 7Li)] | [lithium distribution , lithium concentration gradient, volume variations, defects after ageing] | [M, M, M, Null] | [Ex-situ, operando] | [s, hr] | [operando, ex-situ cell] | [D50-Next] | [Neutron] |  | [Best] | [Commercial] | [Expensive] | [Months] | [Hours/Days] |  |  | [No] |  | [Python]
Imaging | [DTU] | [Laboratory X-ray Tomography, lab XRT] | [Null] | [Null] | [Null] | [Null] | [atomic weight contrast] | [Null] | [Null] | [Null] | [Null] | [Null] | [Null] | [Null] | [Nikon XT] | [laboratory source] |  | [Medium] | [realistic] | [Cheap] | [Days] | [Minutes/Hours]
Imaging | [ILL] | [Laboratory X-ray Tomography, lab XRT] | [Null] | [Null] | [Null] | [Null] | [atomic weight contrast] | [Null] | [Null] | [Null] | [Null] | [Null] | [Null] | [Null] | [D50-NeXT] | [Neutron] |  | [High] | [realistic] | [Cheap] | [Days] | [Minutes/Hours]
Imaging | [CTH] | [Synchrotron X-ray micro-tomography, synchrotron μXRT, μXRT] | [100 x 100 µm2] | [1 µm] | [cm, full cell] | [1 µm] | [few percent in density] | [Null] | [Absorption contrast, phase contrast] | [Morphology] | [Null] | [Ex-situ, operando] | [min] | [operando Swagelok cell] | [ID19, ANATOMIX] | [Synchrotron] | [Yes] | [High] | [realistic] | [Expensive] | [Months] | [Months] | [Medium] | [Medium] | [Yes] | [No] | [Avizo]
Imaging | [CEA] | [Synchrotron X-ray micro-tomography, synchrotron μXRT, μXRT] | [100 x 100 µm2] | [1 µm] | [cm, full cell] | [1 μm] | [few percent in density] | [Null] | [Absorption contrast, phase contrast] | [Morphology] | [H] | [Ex-situ, operando] | [s, min] | [miniaturized version in diameter, Swage-lock type cell] | [ANATOMIX] | [Synchrotron] | [Yes] | [High] | [realistic] | [Expensive] | [Months] | [Hours/Days] | [Medium] | [Medium] | [Yes] | [Yes] | [Inhouse, Python, Opensource, Fiji]
Imaging | [ESRF] | [Synchrotron X-ray micro-tomography, synchrotron μXRT, μXRT] | [100 x 100 µm2] | [1 µm] | [cm, full cell] | [Null] | [few percent in density] | [Null] | [Absorption contrast, phase contrast] | [Morphology] | [Null] | [Ex-situ, operando] | [s, min] | [Null] | [ID19] | [Synchrotron] | [No] | [High] | [realistic] | [Expensive] | [Months] | [Hours/Days] | [Medium] | [Medium] | [Yes] | [Yes] | [Inhouse]
Imaging | [SOLEIL] | [Synchrotron X-ray micro-tomography, synchrotron μXRT, μXRT] | [100 x 100 µm2] | [1 µm] | [cm, full cell] | [Null] | [Null] | [Null] | [Absorption contrast, phase contrast] | [Morphology] | [Null] | [Ex-situ, operando] | [s, min] | [Null] | [PSICHé, ANATOMIX] | [Synchrotron] | [No] | [High] | [realistic] | [Expensive] | [Months] | [Hours/Days] | [Medium] | [Medium] | [Yes] | [Yes] | [Inhouse]
Imaging | [CEA] | [Synchrotron X-ray nano-tomography, synchrotron nXRT, nXRT] | [50 µm x 50 µm] | [20 - 80 nm] | [mm] | [20 - 80 nm] | [few percent in density] | [Null] | [Phase contrast boundary enhancement] | [Morphological information, volume variations, pre-existing defects, after ageing defects] | [H, H, H, H] | [Ex-situ, operando] | [10 min, hrs] | [operando microbattery cell] | [ID16a, ID16b] | [Synchrotron] | [Yes] | [Best] | [realistic] | [Expensive] | [Months] | [Hours/Days] | [Medium] | [Medium/low] | [Yes] | [Yes] | [Inhouse, Python, Opensource, Fiji]
Imaging | [CTH] | [Synchrotron X-ray nano-tomography, synchrotron nXRT, nXRT] | [50 µm x 50 µm] | [20 - 80 nm] | [Full cell] | [20 - 80 nm] | [few percent in density] | [Null] | [Phase contrast boundary enhancement] | [Morphological information, volume variations, pre-existing defects, after ageing defects] | [H, H, H, H] | [Ex-situ, operando] | [10 min, hrs] | [operando cell] | [Null] | [Synchrotron] | [Yes] | [Best] | [realistic] | [Expensive] | [Months] | [Hours/Days] | [Low] | [Low] | [No] | [No] | [Null]
Imaging | [ESRF] | [Synchrotron X-ray nano-tomography, synchrotron nXRT, nXRT] | [50 µm x 50 µm] | [20 - 80 nm] | [Full cell] | [Null] | [Null] | [Null] | [Phase contrast boundary enhancement] | [Morphological information, volume variations, pre-existing defects, after ageing defects] | [H, H, H, H] | [Ex-situ, operando] | [10 min, hrs] | [Null] | [ID16B, ID16A] | [Synchrotron] | [Yes] | [Best] | [realistic] | [Expensive] | [Months] | [Hours/Days] | [Medium] | [Medium/low] | [Yes] | [Yes] | [Inhouse]
Imaging | [SOLEIL] | [Synchrotron X-ray nano-tomography, synchrotron nXRT, nXRT] | [50 µm x 50 µm] | [20 - 80 nm] | [mm, Full cell] | [Null] | [Null] | [Null] | [Phase contrast boundary enhancement] | [Morphological information, volume variations, pre-existing defects, after ageing defects] | [H, H, H, H] | [Ex-situ, operando] | [10 min, hrs] | [Null] | [ANATOMIX] | [Synchrotron] | [No] | [Best] | [realistic] | [Expensive] | [Months] | [Hours/Days] | [Medium] | [Medium/low] | [Yes] | [Yes] | [Inhouse]
Imaging | [ILL] | [Neutron imaging] | [170 x 170 mm3] | [> 4 μm] | [Full cell] | [Null] | [neutron absorption contrast] | [Null] | [neutron absorption contrast, possible isotopes labeling (e.g. 6Li, 7Li) | [lithium distribution , lithium concentration gradient, volume variations, defects after ageing] | [M, M, M, M] | [Ex-situ] | [4 hr] | [Null] | [D50-NeXT] | [Neutron] |  | [High] | [realistic] | [Expensive] | [Months] | [Hours]
Imaging | [ESRF] | [Synchrotron Nano x‑ray fluorescence, nXRF] | [50 μm x 50 μm] | [20 nm, vox. > 10 nm] | [Full cell] | [Null] | [> 0.01 ppm] | [Null] | [Null] | [Elemental analysis] | [H] | [Ex-situ] | [> 1h] | [Null] | [ID16B, ID16A, ID21] | [Synchrotron] | [No] | [Medium] | [realistic] | [Medium] | [Weeks/months] | [Minutes] | [Low] | [Medium/low] | [No] | [Null] | [Inhouse]
Imaging | [CNRS] | [Synchrotron X-ray diffraction computed tomography, synchrotron XRD‑CT] | [5 mm in diameter] | [ 1 µm, vox. > 300 nm] | [full electrode layer (3 mm in diameter)] | [Null] | [crystal contrast, local structure contrast] | [Null] | [crystal contrast, local structure contrast] | [spatially resolved structural information, spatially resolved chemical information, spatially resolved morphological information, diffractogram for each 3D pixel, volume selective analysis of specific battery components, Phase fractions distribution in 3D, Unit cell lattice parameters, atomic site parameters, atomic thermal parameters, lattice strain, lattice size] | [H, H, H, Null],,[Null, Null, H, M, L, L, L] | [Ex-situ, operando] | [min] | [cylindrical cell] | [Null] | [Null] |  | [Low] | [realistic] | [Expensive] | [Months] | [Hours/days]
Imaging | [CEA] | [Synchrotron X-ray diffraction computed tomography, synchrotron XRD‑CT, synchrotron WAXS-CT] | [5 mm in diameter] | [300 nm - 20 µm] | [mm, cm] | [300 nm - 20 µm] | [crystal contrast] | [Null] | [crystal contrast] | [spatially resolved structural information, spatially resolved chemical information, spatially resolved morphological information, diffractogram for each 3D pixel, volume selective analysis of specific battery components, Phase fractions, distribution in 3D, Unit cell lattice parameters, atomic site parameters, atomic thermal parameters, lattice strain, lattice size] | [H, M, M, H, H, H, H, M, L, L, M, M] | [Ex-situ, operando] | [min, hours] | [miniaturized version in diameter, Swage-lock type cell] | [ID31, ID15A] | [Synchrotron] | [Yes] | [High] | [realistic] | [Expensive] | [Months] | [Days/Months] | [Medium] | [Medium] | [Yes] | [Yes] | [Inhouse, Python]
Imaging | [Soleil] | [Synchrotron X-ray diffraction computed tomography, synchrotron XRD‑CT] | [5 mm in diameter] | [ 1 µm, vox. > 300 nm] | [full electrode layer (3 mm in diameter)] | [Null] | [crystal contrast, local structure contrast] | [Null] | [crystal contrast, local structure contrast] | [spatially resolved structural information, spatially resolved chemical information, spatially resolved morphological information, diffractogram for each 3D pixel, volume selective analysis of specific battery components, Phase fractions, distribution in 3D, Unit cell lattice parameters, atomic site parameters, atomic thermal parameters, lattice strain, lattice size] | [H, H, H, Null, Null, Null, H, M, L, L, L] | [Ex-situ, operando] | [min] | [Null] | [PSICHé] | [Synchrotron] |  | [Low] | [realistic] | [Expensive] | [Months] | [Hours/days]
Imaging | [ESRF] | [Synchrotron X-ray diffraction computed tomography, synchrotron XRD‑CT] | [5 mm in diameter] | [ 1 µm, vox. > 300 nm] | [full electrode layer (3 mm in diameter)] | [Null] | [crystal contrast, local structure contrast] | [Null] | [crystal contrast, local structure contrast] | [spatially resolved structural information, spatially resolved chemical information, spatially resolved morphological information, diffractogram for each 3D pixel, volume selective analysis of specific battery components, Phase fractions, distribution in 3D, Unit cell lattice parameters, atomic site parameters, atomic thermal parameters, lattice strain, lattice size] | [H, H, H, Null],,[Null, Null, H, M, L, L, L] | [Ex-situ, operando] | [min] | [Null] | [ID13, ID15A] | [Synchrotron] |  | [Low] | [realistic] | [Expensive] | [Months] | [Hours/days]
Imaging | [CNRS] | [synchrotron pair distribution function computed tomography, synchrotron PDF‑CT] | [5 mm in diameter] | [ 1 µm, vox. > 300 nm] | [full electrode layer (3 mm in diameter)] | [Null] | [crystal contrast, local structure contrast] | [Null] | [crystal contrast, local structure contrast] | [spatially resolved structural information, spatially resolved chemical information, spatially resolved morphological information, diffractogram for each 3D pixel, volume selective analysis of specific battery components, Phase fractions, distribution in 3D, Unit cell lattice parameters, atomic site parameters, atomic thermal parameters, lattice strain, lattice size] | [H, H, H, Null],,[Null, Null, H, M, L, L, L] | [Ex-situ, operando] | [min] | [cylindrical cell] | [Null] | [Null] |  | [Low] | [realistic] | [Expensive] | [Months] | [Hours/days]
Imaging | [Soleil] | [synchrotron pair distribution function computed tomography, synchrotron PDF‑CT] | [5 mm in diameter] | [ 1 µm, vox. > 300 nm] | [full electrode layer (3 mm in diameter)] | [Null] | [crystal contrast, local structure contrast] | [Null] | [crystal contrast, local structure contrast] | [spatially resolved structural information, spatially resolved chemical information, spatially resolved morphological information, diffractogram for each 3D pixel, volume selective analysis of specific battery components, Phase fractions, distribution in 3D, Unit cell lattice parameters, atomic site parameters, atomic thermal parameters, lattice strain, lattice size] | [H, H, H, Null, Null, Null, H, M, L, L, L] | [Ex-situ, operando] | [min] | [Null] | [PSICHé] | [Synchrotron] |  | [Low] | [realistic] | [Expensive] | [Months] | [Hours/days]
Imaging | [ESRF] | [synchrotron pair distribution function computed tomography, synchrotron PDF‑CT] | [5 mm in diameter] | [ 1 µm, vox. > 300 nm] | [full electrode layer (3 mm in diameter)] | [Null] | [crystal contrast, local structure contrast] | [Null] | [crystal contrast, local structure contrast] | [spatially resolved structural information, spatially resolved chemical information, spatially resolved morphological information, diffractogram for each 3D pixel, volume selective analysis of specific battery components, Phase fractions, distribution in 3D, Unit cell lattice parameters, atomic site parameters, atomic thermal parameters, lattice strain, lattice size] | [H, H, H, Null],,[Null, Null, H, M, L, L, L] | [Ex-situ, operando] | [min] | [Null] | [ID13, ID15A] | [Synchrotron] |  | [Low] | [realistic] | [Expensive] | [Months] | [Hours/days]
Imaging | [CEA] | [small angle X-ray scattering computed tomography, SAXS-CT] | [3 mm in diameter] | [10 x 50 x 50 mm3 voxels] | [mm] | [10 x 50 x 50 mm3 voxels] | [Nanostructure contrast] | [Null] | [Nanostructure contrast] | [Short range order, 10 nm to few μm, Pore Structure, Surface area] | [M, M, M, M] | [Ex-situ, operando] | [min, hours] | [Swagelok type cell] | [ID31] | [Synchrotron] | [Yes] | [High] | [realistic] | [Expensive] | [Months] | [Days/Months] | [Medium] | [Medium] | [No] | [Yes] | [Inhouse, Pyhton]
Imaging | [CTH] | [small angle X-ray scattering computed tomography, SAXS-CT] | [3 mm in diameter] | [1 µm] | [Full cell] | [1 µm] | [Null] | [Null] | [Null] | [3D repartition of phases, lithiation mapping] | [Null, Null] | [Ex-situ, in-situ] | [Hours] | [In-situ capillary cell] | [cSAXS] | [Synchrotron] | [Yes] | [High] | [model] | [Expensive] | [Months] | [Months] | [Low] | [Medium] | [No] | [No] | [Matlab]
Imaging | [ESRF] | [small angle X-ray scattering computed tomography, SAXS-CT] | [3 mm in diameter] | [Null] | [5 to 10 mm3 voxels ] | [Null] | [Null] | [Null] | [Null] | [3D repartition of phases, lithiation mapping] | [Null, Null] | [Ex-situ, operando] | [~10 mins per slice] | [Null] | [ID31] | [Synchrotron] |  | [High] | [realistic] | [Medium/Expensive] | [Months] | [Hours]
Imaging | [CTH] | [wide angle X-ray scattering computed tomography, WAXS-CT] | [3 mm in diameter] | [1 µm] | [Full cell] | [1 µm] | [Null] | [Null] | [Null] | [3D repartition of phases, lithiation mapping] | [Null, Null] | [Ex-situ, in-situ] | [Hours] | [In-situ capillary cell] | [cSAXS] | [Synchrotron] | [Yes] | [High] | [model] | [Expensive] | [Months] | [Months] | [Low] | [Medium] | [No] | [No] | [Matlab]
Imaging | [ESRF] | [wide angle X-ray scattering computed tomography, WAXS-CT] | [3 mm in diameter] | [Null] | [5 to 10 mm3 voxels ] | [Null] | [Null] | [Null] | [Null] | [3D repartition of phases, lithiation mapping] | [Null, Null] | [Ex-situ, operando] | [~10 mins per slice] | [Null] | [ID31] | [Synchrotron] |  | [High] | [realistic] | [Medium/Expensive] | [Months] | [Hours]
Imaging | [TUD] | [neutron depth profiling] | [1 cm2] | [Null] | [~40 mm] | [100 nm] | [~1 molar Li density] | [Null] | [Null] | [Li-density as a function of depth] | [M] | [Ex-situ, operando] | [min] | [Pouch type cell] | [NPD] | [Neutron] |  | [Medium] | [realistic] | [Expensive] | [Months] | [Hours]
Imaging | [ESRF] | [Ptychography, Ptychtomography] | [50 mm x 50 mm] | [10 nm, vox.> 5 nm] | [Full cell] | [Null] | [Null] | [Null] | [Phase contrast boundary enhancement] | [Morphological information, volume variations, pre-existing defects, after ageing defects] | [H, H, H, H] | [Ex-situ, operando] | > 10 min | [Null] | [ID16A] | [Synchrotron] |  | [High] | [realistic] | [Expensive] | [Months] | [Hours/days]
Imaging | [SOLEIL] | [Ptychography, Ptychtomography] | 50um x 50um | [10 nm, vox.> 5 nm] | [Full cell] | [Null] | [Null] | [Null] | [Phase contrast boundary enhancement] | [Morphological information, volume variations, pre-existing defects, after ageing defects] | [H, H, H, H] | [Ex-situ, operando] | > 10 min | [Null] | [SWING, HERMES] | [Synchrotron] |  | [High] | [realistic] | [Expensive] | [Months] | [Hours/days]
Imaging | [CNRS] | [Bragg coherent diffraction imaging, BraggCDI] | [Null] | [Null] | [Null] | [Null] | [Null] | [Null] | [Null] | [phase fraction, defects, imaging within grain] | [Null, Null, Null] | [Null] | [Null] | [under development] | [Null] | [Null] |  | [High] | [realistic] | [Expensive] | [Months] | [Hours]
Imaging | [ESRF] | [Bragg coherent diffraction imaging, BraggCDI] | [Null] | [Null] | [Null] | [Null] | [Null] | [Null] | [Null] | [phase fraction, defects, imaging within grain] | [Null, Null, Null] | [Null] | [Null] | [reflection geometry windowed coin cell] | [ID01, ID13, ID10] | [Synchrotron] |  | [High] | [realistic] | [Expensive] | [Months] | [Hours]
Imaging | [SOLEIL] | [Bragg coherent diffraction imaging, BraggCDI] | [Null] | [Null] | [Null] | [Null] | [Null] | [Null] | [Null] | [phase fraction, defects, imaging within grain] | [Null, Null, Null] | [Null] | [Null] | [Null] | [CRISTAL] | [Synchrotron] |  | [High] | [realistic] | [Expensive] | [Months] | [Hours]
Imaging | [CEA] | [Transmission electron microscopy, TEM] | [< 0.1 nm] | [HRTEM: point-to point resolution 0.08 nm] | [transmission, up to 150 nm] | [Null] | [Heavy elements, light elements, Li > 10 at%, > 0.3 at% by EDX] | [EELS: 0.4 - 0.7 eV energy resolution] | [diffraction contrast, thickness contrast, atomic weight contrast] | [atomic structure, crystal structure, chemical composition , morphology, texture, chemical gradients, electronic structure (valence)] | [H, M, M, H, M, M, M] | [Ex-situ] | [Null] | [Ex-situ, post-mortem] | [TITAN] | [Null] |  | [Medium] | [model] | [Cheap] | [Hours/days] | [Hours]
Imaging | [DTU] | [Transmission electron microscopy, TEM] | [< 0.1 nm] | [HRTEM: point-to point resolution 0.08 nm] | [transmission, up to 150 nm] | [Null] | [Heavy elements, light elements, Li > 10 at%, > 0.3 at% by EDX] | [EELS: 0.4 - 0.7 eV energy resolution] | [diffraction contrast, thickness contrast, atomic weight contrast] | [atomic structure, crystal structure, chemical composition , morphology, texture, chemical gradients, electronic structure (valence)] | [H, M, M, H, M, M, M] | [Ex-situ] | [Null] | [In development] | [HRTEM] | [Null] |  | [Medium] | [model] | [Cheap] | [Hours/days] | [Hours]
Imaging | [NIC] | [Transmission electron microscopy, TEM] | [< 0.1 nm] | [HRTEM: point-to point resolution 0.08 nm] | [transmission, up to 150 nm] | [Null] | [Heavy elements, light elements, Li > 10 at%, > 0.3 at% by EDX] | [EELS: 0.4 - 0.7 eV energy resolution] | [diffraction contrast, thickness contrast, atomic weight contrast] | [atomic structure, crystal structure, chemical composition , morphology, texture, chemical gradients, electronic structure (valence)] | [H, M, M, H, M, M, M] | [Ex-situ] | [Null] | [Ex-situ, post-mortem] | [JEM200CF 80-200 kV with QuantumGIF, 100 mm2 Centurio EDX] | [Null] |  | [Medium] | [model] | [Cheap] | [Hours/days] | [Hours]
Imaging | [UOXF] | [Sacnning electron microscopy, SEM] | [probe size ~1 nm, probed area from nm up to hundreds mm] | [SEM ~ 1.5 nm, FIB ~ 1 nm] | [SEM 200 nm to 2 mm, FIB cut up to 20 mm] | [Null] | [EDX 0.3 at%, crystallinity down to nm] | [Null] | [atomic weight] | [surface morphology, chemical composition] | [H, M] | [Ex-situ] | [Null] | [Ex-situ, post-mortem] | [Zeiss Merlin Analytical] | [Null] | [Null] | [Low] | [model] | [Cheap] | [Hours/days] | [Minutes/Hours] | [Low] | [Medium/low] | [No] | [No] | [Null]
Imaging | [CEA] | [Scanning electron microscopy, SEM] | [probe size ~1 nm, probed area from nm up to hundreds mm] | [SEM ~ 1.5 nm, FIB ~ 1 nm] | [SEM 200 nm to 2 mm, FIB cut up to 20 mm] | [Null] | [EDX 0.3 at%, crystallinity down to nm] | [Null] | [atomic weight] | [surface morphology, chemical composition] | [H, M] | [Ex-situ] | [Null] | [Ex-situ, post-mortem] | [FIB-SEM] | [Null] | [Null] | [Low] | [model] | [Cheap] | [Hours/days] | [Minutes/Hours] | [Low] | [Medium/low] | [No] | [No] | [Null]
Imaging | [NIC] | [Scanning electron microscopy, SEM] | [probe size ~1 nm, probed area from nm up to hundreds mm] | [SEM ~ 1.5 nm, FIB ~ 1 nm] | [SEM 200 nm to 2 mm, FIB cut up to 20 mm] | [Null] | [EDX 0.3 at%, crystallinity down to nm] | [Null] | [atomic weight] | [surface morphology, chemical composition] | [H, M] | [Ex-situ] | [Null] | [Ex-situ, post-mortem] | [Zeiss Supra 35 VP, FIB: FEI HeliosNanolab 650] | [Null] | [Null] | [Low] | [model] | [Cheap] | [Hours/days] | [Minutes/Hours] | [Low] | [Medium/low] | [Yes] | [No] | [Inhouse]
Imaging | [ESRF] | [X-ray Laminography] | [Null] | [Null] | [Null] | [Null] | [Null] | [Null] | [Null] | [Null] | [Null] | [Null] | [hr] | [Null] | [ID16B, ID19] | [Synchrotron] |  | [Medium] | [model] | [Expensive] | [Months] | [Minutes/Hours]
Imaging | [ILL] | [Neutron Laminography] | [Null] | [Null] | [Null] | [Null] | [Null] | [Null] | [Null] | [Null] | [Null] | [Null] | [Null] | [Null] | [D50-Next] | [Neutron] |  | [Medium] | [model] | [Expensive] | [Months] | [Minutes/Hours] |  |  | [No]
Imaging | [SOLEIL] | [Scanning transmisison X-ray microscopy, STXM] | [Null] | [Null] | [Null] | [Null] | [Null] | [Null] | [Null] | [Null] | [Null] | [Null] | [Null] | [Null] | [HERMES] | [Synchrotron] | [Yes] | [High] | [model] | [Cheap] | [Days] | [Minutes/Hours] | [Low] | [Medium] | [No] | [Null] | [Commercial]
Imaging | [SOLEIL] | [X-ray photoemission microscopy, XPEEM] | [25 nm] | [Null] | [10 nm] | [0.1 nm] | [Null] | [Null] | [Null] | [SEI chemical imaging, electronic imaging, spectroscopy] | [Null] | [Null] | [100 ms] | [Null] | [HERMES] | [Synchrotron] |  | [High] | [model] | [Medium] | [Days] | [Hours]
Imaging | [SOLEIL] | [Micro X-ray absorption spectroscopy, Micro-XAS] | [Null] | [500 nm] | [10 - 500 mm] | [Null] | [Null] | [Null] | [Null] | [SEI chemical imaging, electronic imaging, spectroscopy] | [Null] | [Null] | [hr] | [Omicron type cell] | [ANTARES] | [Synchrotron] |  | [High] | [model] | [Expensive] | [Months] | [Minutes/Hours]
Imaging | [SOLEIL] | [Nano X-ray photoemission spectroscopy, nano-XPS] | [Null] | [400 nm] | [< 100 nm] | [Null] | [Null] | [Null] | [Null] | [SEI chemical imaging, electronic imaging, spectroscopy] | [Null] | [Null] | [hr] | [Omicron type cell] | [ANTARES] | [Synchrotron] |  | [Best] | [model] | [Expensive] | [Months] | [Minutes/Hours]
Bulk Spectroscopy | [TUD] | [Solid State Nuclear Magnetic Resonance Spectroscopy, NMR] | [Full cell] | [Null] | [~ 10 µm depending on the field and metal] | [Null] | [Depends on element and chemical environment, Convenient nuclei: 1H, 2H, 7Li, 19F, 23Na, 27Al, 31P, 51V] | [Null] | [Null] | [Element specific, chemical environment, Ion mobility (under specific conditions), Convenient nuclei: 1H, 2H, 7Li, 19F, 23Na, 27Al, 31P, 51V, possibility to differentiate different components based on their relaxation time] | [Null, Null, Null, Null, Null] | [Ex-situ, Operando] | [min] | [Capsule cells, Pouch cells] | [500 MHz] | [Null] |  | [High] | [realistic] | [Cheap] | [Days] | [Hours]
Bulk Spectroscopy | [UCAM] | [Solid State Nuclear Magnetic Resonance Spectroscopy, NMR] | [Full cell] | [Null] | [~ 10 µm depending on the field and metal] | [Null] | [Depends on element and chemical environment, Convenient nuclei: 1H, 2H, 7Li, 19F, 23Na, 27Al, 31P, 51V] | [Null] | [Null] | [Element specific, chemical environment, Ion mobility (under specific conditions), Convenient nuclei: 1H, 2H, 7Li, 19F, 23Na, 27Al, 31P, 51V, possibility to differentiate different components based on their relaxation time] | [Null, Null, Null, Null, Null] | [Operando] | [minutes to days] | [Swagelock or coin cells] | [Pfeiffer Omnistar] | [Null] | yes | [High] | [realistic] | [Cheap] | [Days] | [Hours]
Bulk Spectroscopy | [CTH] | [Raman Spectroscopy] | [µm] | [min 0.4 µm] | [hundreds of nm, µm] | [Null] | [Depends on the sample, at least 5 wt%] | [Null] | [Null] | [Electrolyte composition, electrolyte solvation structure, elecrtode structure, intercalation processes ] | [Null, Null, Null, Null] | [Ex-situ, Operando] | [min] | [Operando Raman cell] | [HORIBA LabRam HR Evolution, Bruker FT Raman] | [Null] | [Yes] | [High] | [realistic] | [Medium] | [Weeks] | [Days] | [High] | [High] | [Yes] | [Yes] | [Matlab, Prisma]
Bulk Spectroscopy | [DTU] | [Raman Spectroscopy] | [50 x 50 µm] | [> 1 µm] | [1-5 µm] | [Null] | [Depends on the sample] | [Null] | [Null] | [Electrolyte composition, Electrolyte composition gradients, carbon structure, carbon intercalation processes (anodes)] | [Null, Null, Null, Null] | [Ex-situ, Operando] | [min] | [EL-cell] | [Renishaw] | [Null] |  | [High] | [realistic] | [Cheap] | [Days] | [Hours]
Bulk Spectroscopy | [SOLEIL] | [X-ray Raman Spectroscopy, XRS] | [1 cm x 1 cm] | [20 µm] | [Thin cell, 10 - 100 µm] | [Null] | [Possibility to detect light elements (Lithium)] | [Null] | [Null] | [Chemical state of elemental constituents] | [H] | [Ex-situ] | [> 4 hr] | [Be window (transmission)] | [GALAXIES] | [Synchrotron] |  | [High] | [realistic] | [Expensive] | [Months] | [Hours]
Bulk Spectroscopy | [CEA] | [X-ray Raman Spectroscopy, XRS] | [1 cm x 1 cm] | [20 µm x 2 mm] | [Thin cell, 10 - 100 µm] | [Null] | [Low energy absorption Edges < 600 eV] | [Null] | [Null] | [Chemical state of elemental constituents, oxidation state] | [H, H] | [Ex-situ] | [> 4 hr] | [Ex-situ cell] | [ID20] | [Synchrotron] | [No] | [High] | [realistic] | [Expensive] | [Months] | [Hours] | [High] | [Medium] | [Yes] | [Yes] | [Inhouse, Python]
Bulk Spectroscopy | [ESRF] | [X-ray Raman Spectroscopy, XRS] | [1 cm x 1 cm] | [20 µm] | [Thin cell, 10 - 100 µm] | [Null] | [Possibility to detect light elements (Lithium)] | [Null] | [Null] | [Chemical state of elemental constituents] | [H] | [Ex-situ] | [> 4 hr] | [Null] | [ID20] | [Synchrotron] |  | [High] | [realistic] | [Expensive] | [Months] | [Hours]
Bulk Spectroscopy | [CEA] | [Resonnant inelastic X-ray scattering, RIXS] | [mm] | [20 µm x 20 µm] | [500 nm] | [Null] | [Low energy absorption Edges < 600 eV] | [Null] | [Null] | [Oxidation state, electronic structure] | [H, H] | [Ex-situ] | [> 1 hr] | [Ex-situ] | [ID32] | [Synchrotron] | [No] | [High] | [realistic] | [Expensive] | [Months] | [Hours/Days] | [High] | [Medium] | [Yes] | [Yes] | [Inhouse, Python, opensource, RIXSToolBox]
Bulk Spectroscopy | [Soleil] | [X-ray Absorption Spectroscopy, XAS] | [2.5 µm x 2.5 µm] | [Null] | [Depends of transmission / Fluoresence, Bulk 0.8 to 8 keV] | [Null] | [Null] | [Null] | [Null] | [Oxidation state, 1st and 2nd coordination shell information] | [Null, Null] | [Ex-situ, Operando] | [2 min < Δt < 30 min] | [Null] | [LUCIA] | [Synchrotron] |  | [Best] | [realistic] | [Expensive] | [Months] | [Hours/Days]
Bulk Spectroscopy | [CNRS] | [Hard X-ray Absorption Spectroscopy, Hard-XAS] | [400 µm x 1 mm] | [Null] | [bulk 6.5-43 keV] | [Null] | [> 4 % of element] | [Null] | [Null] | [Mn and Ni redox processes] | [Null] | [Ex-situ, Operando] | [250 ms] | [Kapton windows (transmission), Glassy carbon window [transmission] | [Null] | [Synchrotron] |  | [Best] | [realistic] | [Expensive] | [Months] | [Hours/Days]
Bulk Spectroscopy | [SOLEIL] | [Hard X-ray Absorption Spectroscopy, Hard-XAS] | [400 µm x 1 mm] | [Null] | [bulk 6.5-43 keV] | [Null] | [> 4 % of element] | [Null] | [Null] | [Mn and Ni redox processes] | [Null] | [Ex-situ, Operando] | [250 ms] | [Be window (transmission)] | [ROCK] | [Synchrotron] | [Yes] | [Medium] | [realistic] | [Expensive] | [Months] | [Hours/Days] | [High] | [High] | [Yes] | [Yes/No] | [Inhouse, commercial, opensource]
Bulk Spectroscopy | [SOLEIL] | [Resonant and non resonant Inelastic X-ray Scattering, RIXS / IXS] | [10 - 100 µm] | [Null] | [10 - 100 µm] | [Null] | [Elemental] | [Null] | [Null] | [Oxidation state, chemical environment] | [High / Medium] | [Ex-situ, In-situ, Operando] | [hr] | [Si3N4 window (reflection)] | [GALAXIES] | [Synchrotron] | [Yes] | [Medium] | [realistic] | [Medium] | [Months] | [Hours/Days] | [Null] | [Null] | [Yes] | [Yes] | [opensource]
Bulk Spectroscopy | [CEA] | [Neutron Spectroscopy] | [30 mm x 30 mm] | [Null] | [cm, full cell] | [Null] | [Light elements eg. Li, H, hyperfine splitting of some elements (e.g. Co, Nd, Ho, V,...)] | [Null] | [Null] | [Mechanisms of local diffusion] | [H] | [Ex-situ] | [min to hr] | [Ex-situ] | [IN16, Sharp] | [Neutron] | [No] | [High] | [Model] | [Expensive] | [Months] | [Hours/Days] | [Medium] | [Medium] | [No] | [Yes] | [Inhouse, Python]
Bulk Spectroscopy | [ILL] | [Neutron Spectroscopy] | [30 mm x 30 mm] | [Null] | [cm, full cell] | [Null] | [Light elements eg. Li, H, hyperfine splitting of some elements (e.g. Co, Nd, Ho, V,...)] | [Null] | [Null] | [Mechanisms of local diffusion] | [H] | [Ex-situ] | [min to hr] | [Null] | [IN16B] | [Neutron] |  | [High] | [Model] | [Expensive] | [Months] | [Hours/Days] |  |  | [No]
Bulk Spectroscopy | [CTH] | [Ex-situ Fourier-transform Infrared spectroscopy, FTIR] | [mm] | [Null] | [µm] | [Null] | [Depends on the sample] | [Null] | [Null] | [Interphase composition on metal anode, Electrolyte spectra] | [M, Null] | [Ex-situ] | [mins] | [ex-situ] | [Bruker Alpha] | [Null] | [No] | [High] | [Null] | [Cheap] | [Days] | [Hours] | [High] | [High] | [Yes] | [Yes] | [Matlab, Prisma]
Bulk Spectroscopy | [CSIC] | [Ex-situ Fourier-transform Infrared spectroscopy, FTIR] | [Null] | [Null] | [~ 1 µm] | [Null] | [Depends on the sample] | [Null] | [Null] | [Interphase composition, Electrolyte spectra] | [M, Null] | [Ex-situ, Operando] | [mins] | [Diffuse reflectance cell (HARRIK)] | [Vertex 70 FT-IR Spectrometer] | [Null] |  | [Medium] | [Model] | [Cheap] | [Days] | [Minutes]
Bulk Spectroscopy | [UCAM] | [Electrochemical Mass Spectrometry, EMS] | [Full cell] | [Null] | [bulk] | [Null] | [pmol] | [Null] | [Null] | [Gaseous products from individual chemical events at both electrodes] | [Null, Null, Null, Null, Null] | [Ex-situ, Operando] | [min] | [Capsule cells, Pouch cells] | [200 MHz, 400 MHz, 500 MHz, 700 MHz] | [Null] |  | [High] | [realistic] | [Cheap] | [Days] | [Hours]
Bulk Spectroscopy | [SOLEIL] | [Hard X-Ray Photoelectron Spectroscopy, HAXPES] | [100 µm] | [Null] | [20 nm] | [Null] | [Elemental] | [Null] | [Null] | [Oxidation state, chemical environment] | [High / Medium] | [Ex-situ, In-situ, Operando] | [min to hr] | [Si3N4 window (reflection)] | [GALAXIES] | [Synchrotron] | [Yes] | [Medium] | [realistic] | [Medium] | [Months] | [Hours/Days] | [Null] | [Null] | [Yes] | [No] | [opensource]
Surface spectroscopy | [UU] | [X-ray photoelectron Spectroscopy, XPS] | [ 700 x 300 μm, dia. 110 μm, dia. 55 μm, dia. 27 μm, dia 15 μm, XPS Imaging. 400 x 400 μm, 200 x 200 μm and 80 x 80 μm fields of view. Stitchable (unlimited) to form larger images] | [Imaging: down to < 1 μm] | [ Al Kα: 6-10 nm, Ag Lα:12-20 nm, 'Unlimited' with Gas Cluster Ion Sputtering (GCIS) - depth profiling] | [nm, Gas Cluster Ion Sputtering (GCIS) - nm] | [+/- 5 %] | [Null] | [Null] | [SEI/CEI chemistry (oxidation state), overall chemical composition, imaging: chemical homogemeity of surface] | [H, H, Null] | [Ex-situ, In-situ] | [hrs] | [Various sample holders depending on sample size and technique] | [Kratos AXIS Supra+ (Al Kα, Ag Lα), PHI 5500 (Al Kα)] | [X-ray source] | [No] | [High] | [Model] | [Medium/Expensive] | [Days] | [Minutes/Hours] | [High] | [High] | [Yes] | [Yes] | [Inhouse, commercial, opensource]
Surface spectroscopy | [SOLEIL] | [X-ray photoelectron Spectroscopy, XPS] | [Null] | [Null] | [Null] | [Null] | [+/- 5 %] | [Null] | [Null] | [SEI/CEI chemistry (oxidation state), overall chemical composition, chemical homogemeity of surface] | [H, H, Null] | [Ex-situ, Operando] | [Null] | [omicron sample holder] | [TEMPO] | [Null] |  | [Best] | [Model] | [Medium/Expensive] | [Days] | [Minutes/Hours]
Surface spectroscopy | [CEA] | [X-ray photoelectron Spectroscopy, XPS] | [ 200x200 μm, dia. 100x100 μm, dia. 50x50 μm, dia. 20x20μm, dia 10x10 μm, XPS Imaging.. Stitchable (unlimited) to form larger images] | [Imaging: down to < 7 μm] | [ Al Kα: 6-10 nm, Cr Kα:20-30 nm, 'Unlimited' with Gas Cluster Ion Sputtering (GCIS) - depth profiling] | [nm, Gas Cluster Ion Sputtering (GCIS) - nm] | [+/- 5 %] | [Null] | [Null] | [SEI/CEI chemistry (oxidation state), overall chemical composition, chemical homogemeity of surface] | [H, H, Null] | [Ex-situ, Operando] | [1h to 24h] | [Ulvac-PHI, Omicron sample holder] | [Versa-Probe II Al Kα, Quantes Cr Kα ] | [Null] | [yes/no] | [High] | [Model] | [Medium/Expensive] | [Days] | [Minutes/Hours] | [High] | [High] | [Yes] | [Yes] | [Inhouse, commercial, opensource]
Surface spectroscopy | [NIC] | [X-ray photoelectron Spectroscopy, XPS] | [Null] | [Null] | [Null] | [Null] | [+/- 5 %] | [Null] | [Null] | [SEI/CEI chemistry (oxidation state), overall chemical composition, chemical homogemeity of surface] | [H, H, Null] | [Ex-situ, Operando] | [Null] | [Null] | [retracted] | [Null] |  | [Best] | [Model] | [Medium/Expensive] | [Days] | [Minutes/Hours]
Surface spectroscopy | [SOLEIL] | [X-ray photoelectron Spectroscopy, XPS] | [Null] | [Null] | [Null] | [Null] | [+/- 5 %] | [Null] | [Null] | [SEI/CEI chemistry (oxidation state), overall chemical composition, chemical homogemeity of surface] | [H, H, Null] | [Ex-situ, Operando] | [Null] | [Null] | [TEMPO] | [Synchrotron] |  | [Best] | [Model] | [Medium/Expensive] | [Days] | [Minutes/Hours]
Surface spectroscopy | [ESRF] | [X-ray photoelectron Spectroscopy, XPS] | [Null] | [Null] | [Null] | [Null] | [+/- 5 %] | [Null] | [Null] | [SEI/CEI chemistry (oxidation state), overall chemical composition, chemical homogemeity of surface] | [H, H, Null] | [Ex-situ, Operando] | [Null] | [Null] | [BM25-SPLINE] | [Synchrotron] |  | [Best] | [Model] | [Medium/Expensive] | [Days] | [Minutes/Hours]
Surface spectroscopy | [SOLEIL] | [Nano-X-ray photoelectron Spectroscopy, nano-XPS] | [50 nm] | [Null] | [Up to 4 nm] | [Null] | [+/- 5 %] | [Null] | [Null] | [XPS imaging, chemical homogeneity of surface, electronic homogeneity of surface] | [H, H, Null] | [Null] | [Null] | [Omicron sample holder] | [ANTARES] | [Synchrotron] |  | [Best] | [Model] | [Expensive] | [Months] | [Minutes/Hours]
Surface spectroscopy | [SOLEIL] | [Nano-UPS, Angle resolved photoemisison, nano-ARPES] | [50 nm] | [Null] | [Up to 4 nm] | [Null] | [+/- 5 %] | [Null] | [Null] | [ARPES Imaging, chemical homogeneity of surface, electronic homogeneity of surface] | [H, H, Null] | [Null] | [Null] | [Omicron sample holder] | [ANTARES] | [Synchrotron] |  | [Best] | [Model] | [Expensive] | [Months] | [Minutes/Hours]
Surface spectroscopy | [CNRS] | [(near-) ambient pressure X-ray phtoemission spectroscopy, ((N)AP-XPS)] | [100 x 100 µm] | [Null] | [~ 5 nm at 1.5 keV] | [Null] | [1-10% in composition (depending on cross section)] | [Null] | [Null] | [SEI chemical composition] | [Null] | [Ex-situ, In-situ] | [hrs] | [Cell under development] | [Null] | [Null] |  | [High] | [Model] | [Expensive] | [Months] | [Minutes/Hours]
Surface spectroscopy | [SOLEIL] | [(near-) ambient pressure X-ray phtoemission spectroscopy, ((N)AP-XPS)] | [100 x 100 µm] | [Null] | [~ 5 nm at 1.5 keV] | [Null] | [1-10% in composition (depending on cross section)] | [Null] | [Null] | [SEI chemical composition] | [Null] | [Ex-situ, In-situ] | [hrs] | [TEMPO] | [TEMPO] |  |  | [High] | [Model] | [Expensive] | [Months] | [Minutes/Hours]
Surface spectroscopy | [UU] | [(near-) ambient pressure X-ray phtoemission spectroscopy, ((N)AP-XPS)] | [100 x 100 µm] | [Null] | [~ 5 nm at 1.5 keV] | [Null] | [1-10% in composition (depending on cross section)] | [Null] | [Null] | [SEI chemical composition] | [Null] | [Ex-situ, In-situ] | [hrs] | [HIPPIE (MAXIV)] | [HIPPIE (MAXIV), SPECIES (MAXIV)] | [Synchrotron] | [Yes] | [High] | [Model] | [Expensive] | [Months] | [Minutes/Hours] | [Low] | [Low] | [No] | [Yes] | [Inhouse, commercial, opensource]
Surface spectroscopy | [CNRS] | [hard x-ray photoelectron spectroscopy, HAXPES] | [50 µm] | [Null] | [~30 nm at 10 keV] | [< 1 nm] | [+/- 5%] | [Null] | [Null] | [Local chemical properties, Local electronic properties, SEI chemical composition] | [Null, Null, Null] | [Ex-situ, In-situ] | [mins] | [Cell under development] | [Null] | [Null] |  | [High] | [Model] | [Expensive] | [Months] | [Minutes/Hours]
Surface spectroscopy | [SOLEIL] | [hard x-ray photoelectron spectroscopy, HAXPES] | [50 µm] | [Null] | [~30 nm at 10 keV] | [< 1 nm] | [+/- 5%] | [Null] | [Null] | [Local chemical properties, Local electronic properties, SEI chemical composition] | [Null, Null, Null] | [Ex-situ, In-situ] | [mins] | GALAXIES | [GALAXIES] |  |  | [High] | [Model] | [Expensive] | [Months] | [Minutes/Hours]
Surface spectroscopy | [CEA] | [hard x-ray photoelectron spectroscopy, HAXPES] | [50 µm] | [Null] | [~30 nm at 10 keV] | [< 1 nm] | [+/- 5%] | [Null] | [Null] | [Local chemical properties, Local electronic properties, SEI chemical composition] | [Null, Null, Null] | [Ex-situ, In-situ] | [mins] | [Null] | [In-lab Quantes Cr Kα] | [Null] |  | [High] | [Model] | [Expensive] | [Months] | [Minutes/Hours]
Surface spectroscopy | [UU] | [Reflection electron energy loss spectroscopy, REELS] | [probably 10's - 100's of microns dia.] | [Null] | [~5-10 nm] | [Null] | [Null] | [Null] | [Null] | [band gap, Hydrogen content, conjugation/aromaticity in organics, sp2/sp3 hybridisation in carbon] | [Null] | [Ex-situ, In-situ] | [hrs] | [Various sample holders depending on sample size and technique] | [Kratos AXIS Supra+] | [Electron gun] | [No] | [High] | [Model] | [Medium] | [Days] | [Minutes/Hours] | [Low] | [Low] | [No] | [Yes] | [Null]
Surface spectroscopy | [UU] | [Ion Scattering Spectroscopy, ISS] | [250 - 700 μm dia. Spot size] | [Null] | [surface atomic layer] | [Null] | [ +/- 5 %] | [Null] | [Null] | [chemical composition of surface atomic layer] | [H] | [Ex-situ, Operando] | [Null] | [Various sample holders depending on sample size and technique] | [Kratos AXIS Supra+] | [X-ray source] | [No] | [High] | [Model] | [Medium] | [Days] | [Minutes/Hours] | [Low] | [Low] | [No] | [Yes] | [Null]
Surface spectroscopy | [SOLEIL] | [near edge x-ray absorption fine structure, NEXAFS, Extended x-ray absorption fine structure, EXAFS] | [50 x 50 µm] | [10 µm] | [10 nm, electrode surface] | [Null] | [> 0.01 ppm] | [Null] | [Null] | [Short range order information, coordination number, chemical state of elemental constituents, Chemical structure, Electronic structure] | [H, H, H, Null, Null] | [Ex-situ, Operando] | [2-10 mn] | [omicron sample holder] | [TEMPO] | [Synchrotron] |  | [Medium] | [Model] | [Expensive] | [Months] | [Minutes/Hours]
Surface spectroscopy | [ESRF] | [near edge x-ray absorption fine structure, NEXAFS, Extended x-ray absorption fine structure, EXAFS] | [50 x 50 µm] | [10 µm] | [10 nm, electrode surface] | [Null] | [1 % in composition] | [Null] | [Null] | [Short range order information, coordination number, chemical state of elemental constituents, Chemical structure, Electronic structure] | [H, H, H, Null, Null] | [Ex-situ, Operando] | [2-10 mn] | [Null] | [BM23, ID24, ID26] | [Synchrotron] |  | [Medium] | [Model] | [Expensive] | [Months] | [Minutes/Hours]
Surface spectroscopy | [SOLEIL] | [soft x-ray absorption spectroscopy, XAS] | [100 nm] | [Null] | 10nm | [Null] | [0.01 ppm] | [Null] | [Null] | [Chemical structure, Electronic structure] | [h, h] | [Ex-situ, Operando] | [100 ms – 1 mn] | [Operando electrocell] | [HERMES, ANTARES], | [Synchrotron] |  | [Medium] | [Model] | [Expensive] | [Months] | [Hours]
Surface spectroscopy | [SOLEIL] | [soft X-ray Resonant Inelastic X-ray Scattering, RIXS] | [10 - 100 µm] | [Null] | [< 1 µm] | [Null] | [Null] | [Null] | [Null] | [Null] | [Null] | [Null] | [Null] | [Null] | [SEXTANTS] | [Synchrotron] |  | [High] | [Model] | [Expensive] | [Months] | [Hours/Days]
Surface spectroscopy | [CEA] | [nano-Auger, AES] | [Null] | [20 nm] | [< 5 nm] | [Null] | [Null] | [Null] | [Null] | [Null] | [Null] | [Null] | [Null] | [ULVAC-PHI cell, Omicron cell] | [Auger Phi 700Xi] | [Null] |  | [High] | [Model] | [Cheap] | [Days] | [Minutes]
Surface spectroscopy | [CEA] | [X-ray Reflectometry, XRR] | [1 cm x 1 cm] | [200 µm] | [Thin cell] | [~ 2 nm] | [Null] | [Null] | [Null] | [Roughness, interconnection, phases segregation at the interface, Electronic density at embedded interfaces, Interphase thickness, Interphase composition] | [H, H, H, H, M, L] | [Ex-situ, Operando] | [> 10 min] | cell available, w/o separator (si wafer) | [retracted] | [Null] |  | [High] | [Model] | [Cheap/Medium] | [Days] | [Hours]
Surface spectroscopy | [ESRF] | [X-ray Reflectometry, XRR] | [1 cm x 1 cm] | [200 µm] | [Thin cell] | [~ 2 nm] | [Null] | [Null] | [Null] | [Roughness, interconnection, phases segregation at the interface, Electronic density at embedded interfaces, Interphase thickness, Interphase composition] | [H, H, H, H, M, L] | [Ex-situ, Operando] | [> 10 min] | [Null] | [BM32, ID31] | [Synchrotron] |  | [High] | [Model] | [Cheap/Medium] | [Days] | [Hours]
Surface spectroscopy | [ILL] | [Neutron Reflectometry, NR] | [5 cm x 5 cm] | [5 cm x 5cm] | [~ 500 nm] | [~0.5nm] | [1-10% in composition] | [depends on contrasts - as low as a few %] | [Depends on sample composition - down to 2%] | [Li-density as a function of depth (100 nm resolution, max depth 40 mm), atomic in-plane structure, magnetic in-plane structure, Kinetics of interface evolution of thin solid films, Kinetics of interface evolution of thin multilayers] | [H, H, H, H, M, L] | [Ex-situ, Operando] | [min - hr] | [Null] | [D17] | [Neutron] | [No] | [High] | [Model] | [Expensive] | [Months] | [Hours] | [High, High,High, High, High] | [High, High,High, High, High] | [No] | [No]
Surface spectroscopy | [TUD] | [Neutron Reflectometry, NR] | [2 cm x 5 cm] | [Null] | [~ 500 nm] | [~2nm] | [1-10% in composition] | [Null] | [Null] | [Li-density as a function of depth (100 nm resolution, max depth 40 mm), atomic in-plane structure, magnetic in-plane structure, Kinetics of interface evolution of thin solid films, Kinetics of interface evolution of thin multilayers] | [M, Null, Null, Null, Null] | [Ex-situ, Operando] | [min - hr] | [Null] | [ROG] | [Neutron] |  | [High] | [Model] | [Expensive] | [Months] | [Hours]
Surface spectroscopy | [CTH] | [Grazing incidence small angle X-Ray scattering (GISAXS) and grazing incidence X-Ray difraction (GIXRD), GISAXS/GIXRD] | [mm] | [Null] | [nm] | [Null] | [Null] | [Null] | [Null] | [Interface morphology, interface structure] | [Null, Null] | [Ex-situ, Operando] | [mins] | [Model cell] | [SIXS] | [Synchrotron] | [Yes] | [Medium] | [Model] | [Expensive] | [Months] | [Weeks] | [Low] | [Low] | [Yes] | [No] | [Binoculars, Matlab]